{"text_format":"5","workbook":"xls","n_sheets":1,"sheets":["Sheet1"],"props":{"last_modified_by":"Graeme Gill"},"not_rendered":{"charts":2,"external_links":1},"read":"whole","wrapped":"gz","date_system":"1900","formulas":"cached_only"}

{"sheet":"Sheet1","index":0,"n_rows":257,"n_cols":123,"rows_kept":257,"headers":{"footer":["","Page"]},"flags":["formulas_cached_only"]}
["Gretag i1pro driver vs. ArgyllCMS V1.2.0",{"$nulls":3},"DE is CIE 1976"]
[]
[null,"Overall:",null,"Avg","Max"]
[null,"DE spec Lab:",null,{"v":0.005766884617867666,"d":"0.006"},{"v":0.0116865975373544,"d":"0.012"},"  Difference between Gretag & Argyll spectral to XYZ/Lab"]
[null,"DE to Gretag XYZ.:",null,{"v":0.02648259099006248,"d":"0.026"},{"v":0.048005647063654254,"d":"0.048"},"  Difference between Argyll Measurement and Gretag Lab"]
[null,"DE to Gretag spec.:",null,{"v":0.028020711446945644,"d":"0.028"},{"v":0.050861018186820456,"d":"0.051"},"  Difference between Argyll Measurement and Gretag measurenent from spectral"]
[]
[null,"DE to Argyll Norm:",null,{"v":0.15799510268942302,"d":"0.158"},{"v":0.3534649603921172,"d":"0.353"},"  Difference between Argyll normal and Hi-res measurement"]
[]
["Cal tile"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},85.80172,89.229996,70.703728,95.676453,-0.436507,2.5645379999999998]
["Gretag spec.","DE spec Lab:",{"v":0.008394722032324907,"d":"0.008"},null,85.803892,89.230029,70.695514,95.676467,-0.432508,2.571919,68.204,73.581,77.688,80.477,82.262,83.63,84.494,85.253,86.279,86.947,87.286,87.796,88.204,88.448,88.84,89.082,89.273,89.508,89.348,89.677,89.454,89.435,89.677,89.742,89.718,89.728,89.817,89.478,89.673,90.192,90.242,90.084,90.379,90.603,90.559,90.682]
[null,"DE to Gretag XYZ.:",{"v":0.008764982943505405,"d":"0.009"}]
["Argyll","DE to Gretag spec.:",{"v":0.011934617798658348,"d":"0.012"},null,85.785486,89.211157,70.69036,95.668594,-0.432961,2.562961,68.096,73.643,77.603,80.429,82.216,83.638,84.494,85.265,86.266,86.923,87.291,87.779,88.217,88.432,88.832,89.074,89.245,89.469,89.328,89.644,89.452,89.429,89.646,89.732,89.698,89.694,89.779,89.455,89.644,90.159,90.209,90.058,90.349,90.564,90.509,90.631]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.010046428420091355,"d":"0.010"},null,85.789269,89.214944,70.704473,95.670174,-0.432701,2.5530429999999997,66.262,68.223,70.122,71.968,73.639,75.148,76.5,77.676,78.727,79.676,80.479,81.14,81.735,82.262,82.766,83.248,83.653,83.988,84.273,84.522,84.768,85.023,85.292,85.621,85.963,86.283,86.525,86.752,86.925,87.065,87.18,87.317,87.458,87.626,87.796,87.947,88.088,88.217,88.301,88.365,88.445,88.566,88.7,88.836,88.943,89.016,89.087,89.146,89.192,89.25,89.34,89.419,89.452,89.417,89.367,89.352,89.435,89.553,89.629,89.611,89.543,89.46,89.421,89.409,89.433,89.498,89.571,89.64,89.688,89.719,89.74,89.723,89.713,89.697,89.688,89.693,89.701,89.739,89.773,89.762,89.672,89.559,89.465,89.455,89.519,89.645,89.818,89.986,90.139,90.228,90.235,90.201,90.142,90.089,90.082,90.139,90.227,90.338,90.424,90.501,90.552,90.551,90.52,90.506,90.532,90.565,90.628,90.679,90.741,90.816,90.869,90.896,90.92]
[]
[]
["Dark Skin"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},11.598782,10.063231,4.884697,37.955677,14.255121,15.072365]
["Gretag spec.","DE spec Lab:",{"v":0.003518626152349048,"d":"0.004"},null,11.59938,10.063447,4.884379,37.956062,14.257705,15.074722,5.107,5.718,6.663,7.142,6.883,6.362,5.927,5.65,5.566,5.5760000000000005,5.721,6.019,6.388,6.782,7.32,7.739,8.016,8.296,8.831,9.898,11.286,12.588,13.569,14.338,15.076,15.886,16.929,18.268,20.054,22.024,24.04,26.127,28.473,31.113,34.268,38.294]
[null,"DE to Gretag XYZ.:",{"v":0.027428896295694664,"d":"0.027"}]
["Argyll","DE to Gretag spec.:",{"v":0.027231487840365155,"d":"0.027"},null,11.611418,10.072866,4.89306,37.972889,14.270529,15.057577,5.195,5.75,6.656,7.082,6.848,6.354,5.945,5.666,5.572,5.607,5.741,6.029,6.397,6.799,7.319,7.742,8.021,8.302,8.83,9.906,11.303,12.605,13.586,14.357,15.092,15.899000000000001,16.931,18.287,20.074,22.047,24.066,26.149,28.5,31.136,34.289,38.345]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.061697985137281586,"d":"0.062"},null,11.604155,10.05906,4.887651,37.94822,14.325376,15.043797,4.664,4.933,5.138,5.33,5.649,5.959,6.344,6.682,6.932,7.091,7.172,7.166,7.063,6.937,6.774,6.562,6.383,6.195,6.036,5.923,5.827,5.7059999999999995,5.655,5.5969999999999995,5.582,5.5760000000000005,5.558,5.575,5.578,5.605,5.663,5.718,5.809,5.894,6.023,6.147,6.267,6.381,6.514,6.629,6.767,6.92,7.119,7.311,7.462,7.609,7.739,7.836,7.9350000000000005,8.014,8.088,8.165,8.264,8.398,8.57,8.784,9.076,9.463,9.876,10.326,10.813,11.311,11.789,12.23,12.637,12.98,13.304,13.591,13.854,14.1,14.342,14.591,14.826,15.074,15.332,15.597,15.877,16.181,16.533,16.918,17.31,17.759,18.276,18.896,19.501,20.086,20.708,21.344,22.024,22.737,23.416,24.078,24.713,25.387,26.121,26.898,27.687,28.492,29.3,30.147,31.074,32.072,33.144,34.287,35.527,36.914,38.383,39.869,41.332,42.807,44.204,45.305,45.923]
[]
[]
["Light Skin"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},39.283787,35.508728,20.029224,66.142836,16.60432,16.853445]
["Gretag spec.","DE spec Lab:",{"v":0.006394368225240452,"d":"0.006"},null,39.285114,35.509266,20.026719,66.143251,16.606706,16.859363,10.573,14.08,18.487,20.861,21.271,21.547,22.021,22.819,24.143,25.814,27.364,28.358,29.273,31.189,31.973,30.131,29.32,29.751,29.415,30.078,34.352,41.224,47.251,51.045,53.237,54.689,56.123,57.683,59.7,62,64.802,67.92,71.355,74.501,77.082,79.164]
[null,"DE to Gretag XYZ.:",{"v":0.0264865733910583,"d":"0.026"}]
["Argyll","DE to Gretag spec.:",{"v":0.031795766715078014,"d":"0.032"},null,39.293139,35.522931,20.046164,66.153787,16.586532,16.83716,10.508,14.177,18.547,20.881,21.317,21.556,21.991,22.841,24.175,25.87,27.41,28.385,29.311,31.215,31.992,30.157,29.336,29.767,29.425,30.084,34.355,41.236,47.258,51.048,53.235,54.686,56.126,57.681,59.721,62.034,64.814,67.937,71.379,74.496,77.076,79.188]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.06891690827365643,"d":"0.069"},null,39.261486,35.487329,19.994315,66.126332,16.605302000000002,16.897521,9.471,10.039,10.825,12.21,13.764,15.503,17.132,18.585,19.752,20.52,20.913,21.122,21.234,21.294,21.361,21.417,21.522,21.618,21.734,21.923000000000002,22.168,22.441,22.786,23.176,23.559,24.06,24.587,25.147,25.757,26.381,26.928,27.419,27.822,28.118,28.393,28.582,28.804,29.147,29.696,30.439,31.281,32.019,32.406,32.304,31.712,30.859,30.057,29.505,29.224,29.195,29.361,29.646,29.863,29.885,29.687,29.344,29.129,29.228,29.729,30.691,32.176,34.121,36.381,38.82,41.298,43.602,45.639,47.376,48.864,50.1,51.13,51.984,52.691,53.269,53.777,54.246,54.681,55.126,55.623,56.112,56.549,57.028,57.669,58.484,59.14,59.712,60.352,61.1,61.931,62.858,63.831,64.807,65.774,66.8,67.956,69.135,70.281,71.415,72.473,73.508,74.507,75.469,76.363,77.184,77.907,78.587,79.246,79.819,80.338,80.776,81.123,81.368,81.105]
[]
[]
["Blue Sky"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},17.718979,19.490154,26.983082,51.255801,-5.628205,-21.845169]
["Gretag spec.","DE spec Lab:",{"v":0.004763648286764092,"d":"0.005"},null,17.719121,19.489954,26.980251,51.255571,-5.626451,-21.840746,11.425,17.02,25.453,32.02,34.238,34.711,34.563,34.191,33.459,32.266,30.801,29.095,27.315,25.535,23.588,21.821,20.753,19.967,18.987,17.938,16.919,16.118,15.444,14.802,14.071,13.291,12.608,12.021,11.539,11.158,10.852,10.605,10.462,10.377,10.344,10.364]
[null,"DE to Gretag XYZ.:",{"v":0.03134710535281007,"d":"0.031"}]
["Argyll","DE to Gretag spec.:",{"v":0.03305614177426001,"d":"0.033"},null,17.736401,19.504874,27.010221,51.272728,-5.608027,-21.862168,11.469,17.098,25.52,32.092,34.327,34.747,34.621,34.238,33.48,32.286,30.823,29.118,27.345,25.544,23.587,21.829,20.774,19.985,19.002,17.951,16.94,16.137,15.456,14.82,14.078,13.307,12.627,12.016,11.542,11.153,10.852,10.614,10.477,10.382,10.355,10.384]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.042498334496777986,"d":"0.042"},null,17.734685,19.50392,27.033952,51.271631,-5.612472,-21.904419,9.973,10.76,11.94,13.841,16.103,19.062,22.358,25.558,28.379,30.692,32.38,33.38,34.002,34.41,34.589,34.688,34.745,34.745,34.699,34.591,34.507,34.405,34.264,34.101,33.824,33.535,33.198,32.796,32.347,31.905,31.401,30.898,30.341,29.71,29.118,28.521,27.907,27.324,26.759,26.186,25.582,24.938,24.273,23.641,22.966,22.309,21.785,21.389,21.052,20.767,20.519,20.258,19.986,19.703,19.387,19.014,18.628,18.279,17.93,17.554,17.208,16.921,16.655,16.39,16.128,15.901,15.662,15.466,15.254,15.045,14.834,14.598,14.338,14.069,13.816,13.556000000000001,13.287,13.038,12.808,12.603,12.38,12.161,11.985,11.865,11.708,11.529,11.382,11.237,11.122,11.022,10.943,10.847,10.742,10.665,10.6,10.541,10.513,10.474,10.428,10.383,10.369,10.367,10.366,10.351,10.358,10.352,10.383,10.39,10.406,10.422,10.456,10.515,10.514]
[]
[]
["Foliage"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},10.999722,13.394761,5.553925,43.352131,-13.330124,20.96777]
["Gretag spec.","DE spec Lab:",{"v":0.004763648286764092,"d":"0.005"},null,11.00013,13.394542,5.553151,43.351807,-13.325734,20.97099,4.509,4.958,5.202,5.354,5.47,5.622,5.865,6.268,6.666,6.995,7.255,7.552,8.272,10.632,14.931000000000001,17.888,18.031,16.463,14.459,13.214,12.821,12.104,11.11,10.401,10.32,10.495,10.369,10.099,10.042,10.56,12.377,16.451,22.557,28.545,32.368,34.216]
[null,"DE to Gretag XYZ.:",{"v":0.02633660255233868,"d":"0.026"}]
["Argyll","DE to Gretag spec.:",{"v":0.025952255277717776,"d":"0.026"},null,11.012296,13.409738,5.558868,43.374244,-13.333081,20.981766,4.869,5.108,5.277,5.409,5.443,5.606,5.891,6.274,6.677,7.005,7.254,7.55,8.275,10.656,14.953,17.899,18.045,16.479,14.483,13.232,12.836,12.121,11.129,10.415,10.332,10.499,10.363,10.093,10.048,10.56,12.38,16.481,22.587,28.597,32.386,34.23]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.07702377808183677,"d":"0.077"},null,11.014008,13.421467,5.55632,43.391549,-13.395098,21.02404,4.997,4.959,4.815,4.962,5,5.071,5.186,5.263,5.309,5.349,5.3870000000000005,5.446,5.491,5.515,5.518,5.523,5.5969999999999995,5.667,5.764,5.857,5.997,6.128,6.281,6.441,6.567,6.696,6.802,6.913,7.008,7.091,7.172,7.255,7.351,7.429,7.531,7.672,7.844,8.155,8.627,9.374,10.441,11.812,13.425,15.003,16.344,17.403,18.101,18.429,18.465,18.226,17.772,17.173,16.494,15.826,15.144,14.428,13.777,13.328,13.117,13.073,13.017,12.883,12.703,12.471,12.174,11.825,11.468,11.116,10.797,10.536,10.357,10.268,10.25,10.301,10.411,10.509,10.555,10.531,10.467,10.378,10.264,10.166,10.089,10.064,10.039,10.029,10.084,10.217,10.456,10.87,11.449,12.238,13.278,14.616,16.318,18.284,20.428,22.67,24.892,26.959,28.867,30.427,31.676,32.636,33.357,33.922,34.34,34.7,34.915,35.106,35.303,35.367,35.192]
[]
[]
["Blue Flower"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},25.174962999999998,24.077339,34.470634,56.165145,8.516707,-25.101938]
["Gretag spec.","DE spec Lab:",{"v":0.00525138038995389,"d":"0.005"},null,25.175371,24.07736,34.467208,56.165165,8.518343,-25.096948,11.962,18.42,29.781,40.233,44.211,44.964,44.69,43.848,42.66,41.026,38.925,36.903,34.809,31.933,27.653,23.543,21.341,20.692,20.006,19.307,19.846,21.36,22.831,23.998,24.12,24.284,26.293,31.144,38.761,46.683,52.518,55.731,57.359,58.018,58.327,58.58]
[null,"DE to Gretag XYZ.:",{"v":0.023261045677267462,"d":"0.023"}]
["Argyll","DE to Gretag spec.:",{"v":0.02643215711968914,"d":"0.026"},null,25.194474,24.096836,34.50261,56.184618,8.515308,-25.114584,11.837,18.514,29.898,40.287,44.262,45.04,44.72,43.884,42.698,41.082,38.966,36.924,34.85,31.956,27.68,23.567,21.37,20.707,20.02,19.314,19.858,21.383,22.853,24.008,24.132,24.284,26.306,31.18,38.794,46.727,52.556,55.764,57.397,58.029,58.359,58.6]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.0828480821141919,"d":"0.083"},null,25.17733,24.080925,34.539864,56.168727,8.511281,-25.195794,10.138,11.149,12.535,14.36,17.195,20.823,25.19,29.798000000000002,34.173,37.912,40.719,42.518,43.709,44.469,44.801,44.963,45.031,45.038,44.909,44.728,44.51,44.226,43.929,43.572,43.162,42.769,42.311,41.75,41.211,40.56,39.804,39.008,38.242,37.539,36.93,36.255,35.581,34.865,34.077,33.192,32.131,30.818,29.31,27.772,26.236,24.725,23.481,22.542,21.805,21.251,20.904,20.735,20.691,20.658,20.461,20.05,19.608,19.309,19.175,19.196,19.401,19.781,20.275,20.827,21.395,21.931,22.398,22.83,23.268,23.729,24.12,24.32,24.291,24.117,23.963,23.964,24.148,24.548000000000002,25.176,26.09,27.327,28.955,31.038,33.541,36.175,38.862,41.625,44.344,46.928,49.248,51.21,52.773,54.042,55.061,55.898,56.594,57.093,57.453,57.724,57.866,57.987,58.159,58.304,58.407,58.436,58.459,58.578,58.71,58.762,58.846,58.921,58.945,58.616]
[]
[]
["Blueish Green"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},31.317141,42.663704,34.804672,71.326122,-32.709289,0.5562400000000001]
["Gretag spec.","DE spec Lab:",{"v":0.0064495185091610376,"d":"0.006"},null,31.317324,42.662963,34.800055,71.325616,-32.706436,0.562002,11.679,17.177,25.192,31.492,34.028,35.469,36.924,39.057,42.069,45.594,49.008,52.824,56.662,58.973,59.119,57.253,54.236,49.951,44.792,39.84,35.02,30.407,26.233,23.339,21.782,21.012,20.583,20.311,20.586,21.414,22.546,23.626,24.468,24.668,24.36,24.402]
[null,"DE to Gretag XYZ.:",{"v":0.044928034844627276,"d":"0.045"}]
["Argyll","DE to Gretag spec.:",{"v":0.04836301696544624,"d":"0.048"},null,31.342414,42.687034,34.847875,71.342037,-32.685456,0.521639,11.813,17.234,25.312,31.594,34.094,35.532,36.999,39.106,42.134,45.655,49.042,52.87,56.695,59.011,59.139,57.277,54.264,49.981,44.81,39.862,35.034,30.431,26.249,23.359,21.803,21.039,20.594,20.325,20.612,21.453,22.594,23.659,24.496,24.683,24.367,24.391]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.2694897522263103,"d":"0.269"},null,31.293773,42.701119,34.75063,71.351642,-32.904793,0.6779200000000001,10.558,11.131,12.265,13.978,16.274,19.121,22.238,25.322,28.042,30.211,31.778,32.791,33.523,34.087,34.515,34.97,35.452,35.9,36.38,36.901,37.468,38.143,39.015,39.896,40.823,41.911,43.081,44.303,45.497,46.591,47.72,48.922,50.126,51.438,52.822,54.184,55.541,56.798,57.844,58.651,59.196,59.457,59.497,59.331,58.875,58.163,57.404,56.581,55.587,54.436,53.108,51.626,50.003,48.323,46.576,44.774,43.007,41.371,39.783,38.166,36.541,34.999,33.494,31.973,30.424,28.907,27.473,26.155,24.993,24.021,23.249,22.631,22.131,21.724,21.428,21.201,21.003,20.838,20.706,20.563,20.425,20.297,20.264,20.356,20.434,20.551,20.766,21.051,21.387,21.809,22.209,22.593,22.944,23.314,23.676,24.011,24.309,24.544,24.68,24.727,24.703,24.623,24.503,24.328,24.213,24.159,24.262,24.54,24.912,25.443,26.105,26.777,27.327]
[]
[]
["Orange"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},39.817833,30.687586,4.772728,62.243013,35.086042,57.547328]
["Gretag spec.","DE spec Lab:",{"v":0.005339601951458523,"d":"0.005"},null,39.820113,30.688869,4.7722169999999995,62.244103,35.088449,57.551968,5.719,5.766,5.865,5.742,5.578,5.523,5.469,5.506,5.55,5.655,5.8,5.989,6.326,7.064,8.799,12.092,16.716,22.244,28.732,36.796,44.508,50.208,53.754,55.965,57.473,58.586,59.663,60.534,61.443,62.084,62.623,63.13,63.888,64.465,65.007,65.561]
[null,"DE to Gretag XYZ.:",{"v":0.03354240975243402,"d":"0.034"}]
["Argyll","DE to Gretag spec.:",{"v":0.03684576275503063,"d":"0.037"},null,39.829484,30.705136,4.779522,62.257925,35.058074,57.53635,5.793,5.835,5.873,5.792,5.654,5.534,5.496,5.497,5.549,5.652,5.799,6.006,6.338,7.076,8.811,12.115,16.753,22.282,28.755,36.811,44.526,50.209,53.741,55.974,57.476,58.602,59.664,60.541,61.451,62.067,62.642,63.151,63.911,64.493,65.031,65.582]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.20095559723232867,"d":"0.201"},null,39.86656,30.683279,4.76576,62.239352,35.253641,57.578679,5.907,5.545,5.534,5.714,5.707,5.712,5.776,5.755,5.82,5.823,5.785,5.733,5.623,5.604,5.601,5.543,5.509,5.486,5.502,5.477,5.497,5.495,5.486,5.508,5.517,5.522,5.573,5.608,5.624,5.675,5.7379999999999995,5.808,5.86,5.914,5.983,6.068,6.169,6.284,6.465,6.705,6.996,7.391,7.941,8.627,9.476,10.547,11.838,13.307,14.912,16.63,18.415,20.255,22.199,24.21,26.327,28.662,31.282,34.076,36.888,39.646,42.297,44.75,46.912,48.794,50.399,51.734,52.856,53.818,54.624,55.372,56.038,56.59,57.075,57.508,57.902,58.277,58.639,58.981,59.335,59.684,59.956,60.206,60.566,61.061,61.374,61.481,61.638,61.81,62.008,62.239,62.49,62.658,62.789,62.946,63.164,63.433,63.71,63.951,64.108,64.255,64.459,64.673,64.904,65.076,65.209,65.387,65.566,65.791,65.99,66.152,66.299,66.338,66.005]
[]
[]
["Purpleish Blue"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},12.318707,11.485784,28.297411,40.386897,8.780251,-42.786722]
["Gretag spec.","DE spec Lab:",{"v":0.007023314673284894,"d":"0.007"},null,12.318359,11.485626,28.292955,40.386639,8.778988,-42.779818,10.185,14.858,22.233,28.969,32.587,35.039,37.254,38.779,37.865,34.986,30.942,25.815,21.005,17.255,14.043,11.739,10.635,9.886,9.002,8.367,8.264,8.541,8.73,8.617,8.465,8.685,9.48,10.685,11.962,12.572,12.25,11.661,11.545,11.856,12.477,13.424]
[null,"DE to Gretag XYZ.:",{"v":0.025735282182249963,"d":"0.026"}]
["Argyll","DE to Gretag spec.:",{"v":0.0204866715207714,"d":"0.020"},null,12.321831,11.49101,28.295067,40.395449,8.764669,-42.768111,10.003,14.843,22.197,28.933,32.562,35.034,37.27,38.768,37.875,34.997,30.945,25.815,21.019,17.271,14.052,11.753,10.641,9.889,9.004,8.369,8.268,8.543,8.719,8.625,8.482,8.685,9.484,10.695,11.981,12.57,12.259,11.669,11.546,11.855,12.481,13.444]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.22262026133531146,"d":"0.223"},null,12.333858,11.48477,28.395648,40.385239,8.890598,-42.951407,8.629,9.411,10.575,12.08,14.038,16.513,19.362,22.224,24.849,27.168,29.147,30.612,31.688,32.538,33.258,34.097,34.972,35.74,36.488,37.279,38.1,38.707,39.007,38.994,38.671,38.141,37.307,36.323,35.227,34.048,32.715,31.21,29.533,27.653,25.79,24.034,22.369,20.886,19.568,18.388,17.274,16.162,15.074,14.052,13.124,12.302,11.688,11.254,10.912,10.624,10.384,10.147,9.892,9.628,9.324,8.99,8.687,8.463,8.322,8.232,8.202,8.249,8.328,8.436,8.539,8.653,8.722,8.748,8.731,8.697,8.643,8.569,8.503,8.449,8.455,8.518,8.644,8.843,9.118,9.455,9.832,10.248,10.715,11.223,11.671,12.058,12.355,12.556000000000001,12.642,12.621,12.488,12.296,12.035,11.816,11.631,11.53,11.504,11.508,11.562,11.673,11.818,12.021,12.233,12.477,12.724,13.035,13.396,13.839,14.42,15.102,15.896,16.711,17.428]
[]
[]
["Moderate Red"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},30.269285,20.067474,10.726256,51.913414,47.08931,15.768102]
["Gretag spec.","DE spec Lab:",{"v":0.0044985926688232945,"d":"0.004"},null,30.269982,20.067938,10.725128,51.913936,47.089666,15.772556,9.88,12.356,14.446,14.899000000000001,14.274000000000001,13.78,13.525,13.331,13.108,12.883,12.347,11.436,10.742,10.341,9.748,9.25,9.373,10.089,10.844,11.58,15.464,25.85,39.409,49.986,55.703,58.135,59.133,59.424,59.597,59.472,59.32,59.135,59.209,59.125,59.083,59.065]
[null,"DE to Gretag XYZ.:",{"v":0.01910621545465978,"d":"0.019"}]
["Argyll","DE to Gretag spec.:",{"v":0.01570927372605182,"d":"0.016"},null,30.275484,20.074018,10.725416,51.920795,47.080692,15.783474,9.699,12.346,14.507,14.883,14.306,13.791,13.541,13.339,13.09,12.865,12.345,11.434,10.747,10.344,9.745,9.257,9.375,10.089,10.841,11.594,15.488,25.887,39.435,49.987,55.697,58.121,59.112,59.398,59.571,59.439,59.304,59.112,59.188,59.126,59.088,59.071]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.2446001685404194,"d":"0.245"},null,30.261853,20.025057,10.74198,51.86553,47.267892,15.636059,9.086,9.5,10.271,11.23,12.139,13.252,14.117,14.674,15.006,15.063,15.016,14.826,14.571,14.381,14.148,13.94,13.782,13.671,13.582,13.562,13.506,13.441,13.347,13.273,13.182,13.087,13.038,12.948,12.913,12.819,12.662,12.429,12.113,11.752,11.416,11.124,10.879,10.711,10.594,10.507,10.389,10.212,10.006,9.764,9.527,9.317,9.204,9.176,9.218,9.323,9.496,9.747,10.075,10.42,10.706,10.902,11.026,11.128,11.367,11.91,13.019,14.895,17.613,21.22,25.559,30.273,35.031,39.585,43.763,47.392,50.365,52.721,54.576,55.963,56.988,57.713,58.24,58.621,58.917,59.13,59.242,59.277,59.391,59.657,59.718,59.599,59.505,59.442,59.391,59.379,59.374,59.318,59.195,59.11,59.102,59.176,59.2,59.207,59.189,59.133,59.085,59.12,59.133,59.116,59.081,59.044,59.053,59.107,59.077,59.058,59.093,59.094,58.722]
[]
[]
["Purple"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},8.265759,6.391304,10.743704,30.378976,20.556743,-21.415147]
["Gretag spec.","DE spec Lab:",{"v":0.004161553916508229,"d":"0.004"},null,8.26627,6.391342,10.743822,30.379067,20.560898,-21.415362,8.458,12.078,16.398,19.225,19.715,18.656,16.646,14.304,12.004,10.043,8.614,7.515,6.607,6.026,5.6850000000000005,5.378,5.18,5.183,5.2780000000000005,5.299,5.156,5.174,5.688,6.987,9.128,11.59,13.858,15.862,18,20.353,23.077,26.295,29.828,33.487,37.54,42.52]
[null,"DE to Gretag XYZ.:",{"v":0.03906358315618082,"d":"0.039"}]
["Argyll","DE to Gretag spec.:",{"v":0.03679056979716512,"d":"0.037"},null,8.280056,6.401464,10.764107,30.403538,20.57791,-21.436933,8.72,12.063,16.352,19.299,19.756,18.678,16.669,14.336,12.046,10.055,8.619,7.522,6.621,6.032,5.7010000000000005,5.382,5.184,5.181,5.287,5.317,5.171,5.191,5.694,6.997,9.142,11.605,13.87,15.891,18.012,20.362,23.102,26.327,29.855,33.507,37.587,42.545]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.1724317606794055,"d":"0.172"},null,8.273922,6.392424,10.805123,30.381683,20.617631,-21.603298,7.732,8.191,8.965,10.213,11.652,13.227,14.922,16.433,17.788,18.775,19.432,19.821,19.936,19.911,19.74,19.356,18.81,18.209,17.466,16.707,15.928,15.148,14.39,13.615,12.846,12.077,11.384,10.72,10.089,9.538,9.068,8.647,8.236,7.862,7.519,7.194,6.871,6.575,6.341,6.171,6.027,5.901,5.798,5.718,5.618,5.492,5.383,5.299,5.224,5.169,5.145,5.146,5.163,5.203,5.252,5.295,5.333,5.351,5.329,5.2780000000000005,5.22,5.158,5.121,5.12,5.145,5.237,5.395,5.619,5.941,6.366,6.911,7.558,8.296,9.112,9.972,10.824,11.648,12.441,13.209,13.927,14.594,15.229,15.906,16.623,17.34,18.031,18.745,19.518,20.333,21.199,22.124,23.072,24.087,25.163,26.314,27.523,28.714,29.891,31.038,32.216,33.455,34.807,36.186,37.636,39.172,40.82,42.584,44.438,46.326,48.296,50.225,51.875,52.939]
[]
[]
["Yellow Green"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},35.544086,44.359955,8.785229,72.468437,-22.817703,57.730569]
["Gretag spec.","DE spec Lab:",{"v":0.006464832093720342,"d":"0.006"},null,35.545431,44.359507,8.784297,72.46814,-22.811902,57.733407,5.911,6.108,6.34,6.351,6.42,6.634,6.942,7.492,8.532,10.263,13.057,17.603,25.25,36.283,47.429,53.939,55.933,55.146,52.728,49.845,46.269,42.358,38.305,35.236,33.502,32.617,32.113,31.832,32.235,33.295,34.683,35.893,36.852,37.064,36.723,36.808]
[null,"DE to Gretag XYZ.:",{"v":0.026560931346628087,"d":"0.027"}]
["Argyll","DE to Gretag spec.:",{"v":0.03216733078762058,"d":"0.032"},null,35.541933,44.362156,8.79227,72.4699,-22.831251,57.70777,5.841,6.062,6.313,6.423,6.489,6.611,6.921,7.511,8.533,10.282,13.073,17.635,25.276,36.316,47.472,53.956,55.944,55.15,52.709,49.813,46.263,42.342,38.3,35.235,33.501,32.613,32.118,31.832,32.248,33.315,34.694,35.917,36.854,37.069,36.742,36.835]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.3534649603921172,"d":"0.353"},null,35.536693,44.403571,8.715298,72.497423,-22.967501,58.032756,5.89,5.857,5.9719999999999995,5.997,6.038,6.149,6.155,6.27,6.357,6.457,6.493,6.424,6.41,6.446,6.483,6.523,6.61,6.7,6.795,6.908,7.063,7.224,7.438,7.717,8.041,8.419,8.891,9.428,10.084,10.869,11.787,12.86,14.096,15.618,17.365,19.436,21.956,24.961,28.404,32.25,36.257,40.342,44.299,47.696,50.423,52.574,54.19,55.282,55.947,56.197,56.138,55.79,55.251,54.572,53.738,52.777,51.802,50.845,49.837,48.691,47.485,46.301,45.089,43.78,42.406,40.981,39.589,38.241,37.02,35.977,35.141,34.441,33.864,33.431,33.1,32.83,32.602,32.414,32.243,32.105,31.935,31.791,31.764,31.917,32.051,32.196,32.488,32.839,33.265,33.748,34.256,34.714,35.135,35.537,35.951,36.349,36.672,36.927,37.065,37.119,37.105,37.025,36.894,36.739,36.601,36.605,36.744,37.043,37.486,38.125,38.829,39.493,39.887]
[]
[]
["Orange Yellow"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},49.347069,43.774529,6.193294,72.077534,20.301509,67.484826]
["Gretag spec.","DE spec Lab:",{"v":0.005833842130190457,"d":"0.006"},null,49.349615,43.775434,6.192657,72.078141,20.305772,67.488762,5.848,5.934,6.111,6.153,6.148,6.141,6.167,6.279,6.6,7.364,8.632,10.155,11.703,13.834,18.085,25.55,35.125,44.205,50.58,55.098,58.04,60.213,61.862,63.174,64.235,65.143,66.092,66.889,67.758,68.288,68.78,69.205,69.892,70.398,70.867,71.407]
[null,"DE to Gretag XYZ.:",{"v":0.017492451114694156,"d":"0.017"}]
["Argyll","DE to Gretag spec.:",{"v":0.022910283149708572,"d":"0.023"},null,49.350857,43.781291,6.19812,72.082069,20.292195,67.470731,5.893,5.944,6.088,6.143,6.107,6.105,6.187,6.304,6.625,7.359,8.629,10.166,11.713,13.85,18.088,25.569,35.151,44.221,50.586,55.099,58.024,60.205,61.855,63.175,64.239,65.153,66.096,66.891,67.74,68.283,68.786,69.221,69.895,70.425,70.907,71.418]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.2466958717226514,"d":"0.247"},null,49.354072,43.731201,6.14362,72.048465,20.445725,67.660884,6.265,6.001,5.796,5.838,5.947,5.963,6.061,6.188,6.176,6.15,6.143,6.103,6.066,6.076,6.083,6.047,6.067,6.096,6.119,6.126,6.157,6.205,6.266,6.328,6.422,6.546,6.715,6.948,7.243,7.622,8.065,8.554,9.078,9.624,10.154,10.637,11.131,11.647,12.214,12.877,13.665,14.692,16.027,17.685,19.675,22.114,25.016,28.198,31.571,35.047,38.399,41.537,44.419,46.893,48.969,50.78,52.436,53.928,55.191,56.252,57.216,58.097,58.887,59.609,60.254,60.824,61.348,61.828,62.278,62.725,63.156,63.564,63.902,64.228,64.547,64.858,65.153,65.439,65.774,66.07,66.307,66.538,66.881,67.387,67.68,67.768,67.874,68.035,68.211,68.403,68.603,68.775,68.865,68.993,69.237,69.487,69.678,69.878,70.045,70.194,70.346,70.569,70.792,70.932,71.05,71.215,71.413,71.639,71.785,71.9,72.051,72.13,71.631]
[]
[]
["Blue"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},7.425426,6.044739,23.65612,29.525059,16.497714,-53.398175]
["Gretag spec.","DE spec Lab:",{"v":0.0116865975373544,"d":"0.012"},null,7.424596,6.044549,23.650247,29.524582,16.491842,-53.388082,7.174,9.547,13.31,17.718,21.303,24.921,29.534,34.477,35.616,32.147,26.013,19.04,13.197,9.194,6.795,5.519,4.879,4.473,4.201,4.045,3.966,3.912,3.8890000000000002,3.88,3.911,3.968,4.042,4.126,4.229,4.299,4.336,4.385,4.515,4.775,5.134,5.681]
[null,"DE to Gretag XYZ.:",{"v":0.020817714307770473,"d":"0.021"}]
["Argyll","DE to Gretag spec.:",{"v":0.02962865509941539,"d":"0.030"},null,7.433201,6.051259,23.67794,29.541421,16.501413,-53.410503,6.89,9.477,13.37,17.766,21.361,25.003,29.556,34.504,35.647,32.183,26.039,19.057,13.211,9.215,6.803,5.53,4.881,4.478,4.202,4.051,3.963,3.919,3.8970000000000002,3.891,3.917,3.967,4.047,4.122,4.222,4.303,4.349,4.395,4.53,4.779,5.141,5.711]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.33030889829673216,"d":"0.330"},null,7.448807,6.044322,23.7924,29.524012,16.725270000000002,-53.652761,6.263,6.785,7.261,8.053,9.082,10.378,11.735,13.213,14.775,16.26,17.675,19.071,20.201,21.107,22.138,23.385,24.777,26.137,27.603,29.384,31.324,33.283,34.823,35.825,36.25,36.084,35.393,34.215,32.677,30.813,28.639,26.296,23.87,21.342,18.988,16.822,14.809,13.033,11.508,10.214,9.115,8.168,7.385,6.775,6.266,5.833,5.5,5.259,5.045,4.866,4.723,4.59,4.472,4.366,4.276,4.194,4.123,4.078,4.044,4.002,3.973,3.955,3.946,3.934,3.909,3.899,3.899,3.898,3.8890000000000002,3.8810000000000002,3.8890000000000002,3.895,3.908,3.9210000000000003,3.931,3.952,3.974,3.991,4.02,4.058,4.075,4.095,4.126,4.173,4.216,4.234,4.255,4.279,4.308,4.318,4.33,4.337,4.351,4.359,4.3870000000000005,4.433,4.473,4.524,4.594,4.68,4.779,4.896,5.013,5.156,5.295,5.468,5.666,5.926,6.2620000000000005,6.667,7.15,7.69,8.194]
[]
[]
["Green"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},15.27335,23.72168,7.658977,55.80805,-38.981045,33.243493]
["Gretag spec.","DE spec Lab:",{"v":0.005390442746194598,"d":"0.005"},null,15.273634,23.720954,7.658081,55.807318,-38.97621,33.245761,5.555,5.978,6.231,6.354,6.37,6.49,6.781,7.246,8.027,9.386,11.386,14.127,18.06,24.178,31.836,36.081,36.048,33.311,29.147,24.853,20.665,16.751,13.176,10.566,9.094,8.37,7.956,7.615,7.434,7.445,7.6530000000000005,7.976,8.425,8.805,8.887,8.806000000000001]
[null,"DE to Gretag XYZ.:",{"v":0.03096520025770902,"d":"0.031"}]
["Argyll","DE to Gretag spec.:",{"v":0.03139313960087366,"d":"0.031"},null,15.277531,23.729979,7.655652,55.816423,-38.992452,33.271036,5.616,5.941,6.175,6.28,6.353,6.505,6.759,7.231,8.033,9.385,11.392,14.142,18.075,24.204,31.86,36.103,36.057,33.317,29.148,24.854,20.672,16.753,13.177,10.572,9.112,8.381,7.956,7.616,7.431,7.439,7.652,7.989,8.447,8.812,8.87,8.798]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.3081553972008221,"d":"0.308"},null,15.252706,23.742706,7.602446,55.829261,-39.19441,33.503432,5.409,5.5440000000000005,5.644,5.709,5.808,5.995,6.111,6.218,6.239,6.23,6.283,6.306,6.343,6.391,6.385,6.431,6.517,6.574,6.647,6.734,6.854,7.03,7.192,7.389,7.666,7.94,8.311,8.721,9.233,9.831,10.511,11.269,12.084,13.023,14.036,15.142,16.416,17.861,19.583,21.603,23.963,26.652,29.485,32.03,34.031,35.474,36.38,36.751,36.709,36.301,35.571,34.579,33.4,32.079,30.653,29.127,27.608,26.182,24.798000000000002,23.378,21.977,20.636,19.348,18.054,16.755,15.464,14.247,13.117,12.079,11.196,10.471,9.887,9.411,9.049,8.76,8.542,8.36,8.203,8.065,7.949,7.82,7.686,7.59,7.525,7.468,7.407,7.392,7.396,7.426,7.47,7.55,7.64,7.737,7.845,7.97,8.128,8.293,8.456,8.598,8.734,8.819,8.891,8.907,8.893,8.864,8.808,8.782,8.809,8.856,8.962,9.116,9.265,9.412]
[]
[]
["Red"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},22.617523,12.839955,4.051586,42.521102,56.117143,27.654589]
["Gretag spec.","DE spec Lab:",{"v":0.003694680094407115,"d":"0.004"},null,22.618155,12.840185,4.051119,42.52145,56.118512,27.658003,4.994,5.053,5.162,5.126,5.098,5.003,4.992,5.015,4.975,4.831,4.715,4.614,4.55,4.469,4.454,4.475,4.549,4.674,4.93,5.55,7.016,10.868,19.158,32.468,47.923,59.936,66.581,69.608,71.229,71.988,72.547,72.948,73.493,73.757,73.948,74.216]
[null,"DE to Gretag XYZ.:",{"v":0.048005647063654254,"d":"0.048"}]
["Argyll","DE to Gretag spec.:",{"v":0.050861018186820456,"d":"0.051"},null,22.640656,12.858253,4.065575,42.548887,56.102473,27.618294,5.197,5.184,5.271,5.198,5.101,5.04,5.032,5.025,4.967,4.851,4.724,4.638,4.5600000000000005,4.493,4.468,4.496,4.559,4.675,4.938,5.563,7.038,10.887,19.195,32.527,47.964,59.957,66.591,69.606,71.237,72.002,72.538,72.917,73.455,73.763,73.97,74.214]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.2406136078882471,"d":"0.241"},null,22.579142,12.796202,4.068679,42.454553,56.229461,27.436993,5.451,5.287,5.2780000000000005,5.209,5.224,5.266,5.266,5.212,5.182,5.174,5.165,5.156,5.14,5.119,5.1,5.071,5.042,5.045,5.047,5.007,5.006,5.042,5.031,5.046,5.018,4.968,4.938,4.897,4.856,4.803,4.77,4.73,4.703,4.659,4.6370000000000005,4.598,4.572,4.554,4.522,4.513,4.4879999999999995,4.478,4.461,4.465,4.468,4.47,4.489,4.513,4.534,4.548,4.577,4.609,4.66,4.725,4.806,4.914,5.057,5.254,5.502,5.82,6.256,6.864,7.696,8.868,10.49,12.631,15.331,18.644,22.601,27.181,32.222,37.578,43.047,48.324,53.088,57.135,60.482,63.162,65.281,66.882,68.027,68.9,69.715,70.551,71.064,71.307,71.522,71.73,71.946,72.164,72.395,72.542,72.641,72.74,72.933,73.142,73.305,73.465,73.59,73.619,73.709,73.833,73.952,74.015,74.014,74.082,74.193,74.337,74.409,74.508,74.632,74.68,74.159]
[]
[]
["Yellow"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},58.738396,59.294525,7.213908,81.453219,3.802263,79.248661]
["Gretag spec.","DE spec Lab:",{"v":0.006104062991158751,"d":"0.006"},null,58.741115,59.294962,7.21337,81.453458,3.807771,79.251281,5.74,5.386,5.414,5.455,5.457,5.42,5.51,5.675,6.106,7.259,10.033,15.823,24.995,35.444,45.39,53.559,59.298,62.992,65.319,67.466,68.911,70.177,71.198,72.116,72.897,73.581,74.345,74.979,75.691,76.069,76.399,76.674,77.191,77.55,77.887,78.261]
[null,"DE to Gretag XYZ.:",{"v":0.011581236073930298,"d":"0.012"}]
["Argyll","DE to Gretag spec.:",{"v":0.01569418401830281,"d":"0.016"},null,58.740069,59.296832,7.217225,81.454482,3.800841,79.237237,5.51,5.298,5.439,5.441,5.401,5.433,5.519,5.689,6.095,7.259,10.052,15.847,25.028,35.456,45.403,53.571,59.302,62.998,65.313,67.445,68.906,70.16,71.202,72.124,72.917,73.578,74.343,74.967,75.675,76.073,76.398,76.642,77.164,77.554,77.86,78.251]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.306733041410617,"d":"0.307"},null,58.764132,59.318837,7.147789,81.466536,3.806757,79.543676,5.909,5.34,5.396,5.312,5.264,5.324,5.359,5.438,5.441,5.393,5.37,5.36,5.415,5.454,5.418,5.423,5.442,5.418,5.455,5.529,5.573,5.607,5.678,5.761,5.858,6.038,6.271,6.583,7.052,7.682,8.547,9.72,11.202,13.157,15.521,18.279,21.486,24.956,28.516,32.06,35.482,38.865,42.233,45.365,48.246,50.974,53.529,55.792,57.77,59.459,60.866,62.035,63.074,63.965,64.72,65.408,66.13,66.863,67.507,68.024,68.488,68.945,69.409,69.836,70.235,70.58,70.9,71.208,71.516,71.832,72.148,72.452,72.71,72.95,73.162,73.388,73.61,73.834,74.1,74.368,74.52,74.668,74.995,75.472,75.699,75.729,75.816,75.904,76.021,76.174,76.318,76.413,76.461,76.511,76.689,76.883,77.034,77.191,77.291,77.4,77.521,77.68,77.829,77.909,77.989,78.099,78.258,78.411,78.543,78.681,78.765,78.711,78.271]
[]
[]
["Magenta"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},30.030071,19.714593999999998,22.72579,51.512978,47.917774,-13.735843]
["Gretag spec.","DE spec Lab:",{"v":0.003784951386738705,"d":"0.004"},null,30.031121,19.714917,22.724586,51.513347,47.920134,-13.732907,11.612,17.365,26.461,33.257,35.113,34.521,32.681,30.046,27.17,24.359,21.59,18.913,16.834,15.015,12.812,10.966,10.427,10.712,10.739,11.233,13.695,18.648,26.004,36.372,48.978,60.496,68.734,73.763,77.029,78.734,79.568,79.994,80.633,81.128,81.673,82.427]
[null,"DE to Gretag XYZ.:",{"v":0.03172582268437205,"d":"0.032"}]
["Argyll","DE to Gretag spec.:",{"v":0.03348936458041412,"d":"0.033"},null,30.054926,19.732602,22.756766,51.533529,47.922674,-13.759511,11.655,17.376,26.438,33.255,35.097,34.585,32.713,30.098,27.227,24.391,21.613,18.954,16.857,15.035,12.829,10.965,10.438,10.721,10.758,11.245,13.707,18.667,26.034,36.4,49.008,60.523,68.77,73.806,77.046,78.732,79.598,80.033,80.69,81.153,81.673,82.436]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.28003021850150445,"d":"0.280"},null,30.013546,19.675718,22.825196,51.468572,48.047001,-14.001875,10.27,11.09,12.285,14.027,16.518,19.572,22.941,26.486,29.613,31.975,33.513,34.401,34.996,35.31,35.336,35.162,34.787,34.231,33.513,32.781,31.97,31.096,30.159,29.216,28.236,27.311,26.395,25.449,24.496,23.605,22.656,21.7,20.743,19.77,18.91,18.139,17.426,16.81,16.257,15.692,15.106,14.418,13.65,12.852,12.094,11.397,10.868,10.546,10.38,10.361,10.466,10.616,10.754,10.834,10.809,10.734,10.693,10.793,11.078,11.591,12.402,13.512,14.897,16.546,18.464,20.622,23.008,25.684,28.756,32.268,36.131,40.302,44.687,49.132,53.398,57.338,60.885,64.023,66.777,69.08,70.942,72.52,73.996,75.458,76.494,77.195,77.777,78.295,78.738,79.134,79.422,79.627,79.711,79.834,80.024,80.273,80.488,80.699,80.821,80.939,81.095,81.276,81.506,81.706,81.893,82.12,82.451,82.767,83.054,83.334,83.618,83.747,83.358]
[]
[]
["Cyan"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},13.412623,19.248283,30.16923,50.976429,-29.620464,-27.550268]
["Gretag spec.","DE spec Lab:",{"v":0.007642006150219448,"d":"0.008"},null,13.411909,19.247829,30.16406,50.975902,-29.622791,-27.543008,9.935,13.795,19.325,24.107,26.491,28.646,31.479,34.994,38.834,42.68,44.795,44.917,43.082,39.416,34.429,28.876,23.778,18.957,14.725999999999999,11.709,9.836,8.737,7.965,7.451,7.218,7.142,7.159,7.271,7.4879999999999995,7.643,7.5809999999999995,7.399,7.174,7.065,7.2620000000000005,7.865]
[null,"DE to Gretag XYZ.:",{"v":0.026990069544184418,"d":"0.027"}]
["Argyll","DE to Gretag spec.:",{"v":0.02247192839967477,"d":"0.022"},null,13.421833,19.262805,30.177241,50.993268,-29.633762,-27.533895,9.795,13.874,19.407,24.091,26.465,28.641,31.485,35.025,38.851,42.706,44.82,44.945,43.088,39.451,34.455,28.897,23.794,18.97,14.746,11.724,9.84,8.748,7.975,7.455,7.224,7.148,7.165,7.281,7.499,7.642,7.5969999999999995,7.4030000000000005,7.181,7.058,7.266,7.878]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.25068031410942515,"d":"0.251"},null,13.393766,19.274151,30.141411,51.006419,-29.871202,-27.454586,8.776,9.207,10.289,11.588,13.153,15.183,17.36,19.417,21.277,22.914,24.272,25.202,25.799,26.329,27.029,27.853,28.518,29.242,30.272,31.435,32.655,33.829,34.989,36.151,37.332,38.634,40.058,41.447,42.705,43.705,44.453,44.984,45.233,45.258,45.113,44.723,44.044,43.216,42.196,40.98,39.595,38.048,36.355,34.648,32.82,30.884,29.021,27.281,25.578,23.856,22.194,20.539,18.917,17.389,15.957,14.619,13.435,12.439,11.618,10.894,10.264,9.777,9.379,9.036,8.732,8.465,8.201,7.963,7.746,7.566,7.423,7.337,7.264,7.216,7.187,7.168,7.15,7.149,7.144,7.153,7.19,7.225,7.276,7.362,7.447,7.522,7.594,7.631,7.654,7.654,7.642,7.605,7.552,7.47,7.404,7.347,7.255,7.158,7.087,7.032,7.02,7.074,7.149,7.2620000000000005,7.415,7.586,7.829,8.163,8.551,9.079,9.693,10.388,11.101]
[]
[]
["White"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},85.792603,89.321831,70.295708,95.714753,-0.618624,2.996727]
["Gretag spec.","DE spec Lab:",{"v":0.008172265964834923,"d":"0.008"},null,85.79489,89.321867,70.287946,95.714768,-0.6144149999999999,3.003732,14.809,22.568,40.024,64.348,79.036,83.188,84.436,85.453,86.408,87.003,87.443,87.938,88.372,88.613,88.935,89.13,89.417,89.548,89.377,89.72,89.621,89.643,89.689,89.735,89.769,89.793,89.965,90.024,90.239,90.164,90.083,89.937,90.153,90.206,90.211,90.325]
[null,"DE to Gretag XYZ.:",{"v":0.022168170898836534,"d":"0.022"}]
["Argyll","DE to Gretag spec.:",{"v":0.01589437051286195,"d":"0.016"},null,85.769995,89.293457,70.254588,95.702922,-0.6098779999999999,3.013309,14.815,22.631,39.96,64.212,78.941,83.105,84.427,85.445,86.359,86.957,87.387,87.91,88.328,88.589,88.928,89.105,89.375,89.499,89.339,89.684,89.599,89.641,89.664,89.721,89.775,89.749,89.921,90.006,90.226,90.138,90.057,89.929,90.115,90.173,90.175,90.3]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.02796297983048163,"d":"0.028"},null,85.749215,89.263828,70.253721,95.690566,-0.595455,2.992785,13.647,13.977,15.32,17.582,20.971,25.642,31.6,38.879,47.31,56.133,64.511,71.412,76.27,79.427,81.32,82.437,83.201,83.705,84.057,84.383,84.746,85.107,85.477,85.792,86.091,86.329,86.57,86.748,86.911,87.062,87.198,87.363,87.549,87.704,87.899,88.052,88.178,88.268,88.4,88.509,88.576,88.681,88.772,88.908,88.96,88.936,89.009,89.16,89.304,89.392,89.432,89.443,89.426,89.4,89.356,89.319,89.404,89.541,89.65,89.649,89.58,89.571,89.61,89.646,89.652,89.636,89.61,89.617,89.603,89.636,89.68,89.729,89.756,89.753,89.742,89.736,89.744,89.772,89.839,89.881,89.854,89.82,90.002,90.368,90.412,90.237,90.119,90.045,90.026,90.069,90.076,90.06,89.988,89.933,89.983,90.019,90.052,90.111,90.131,90.121,90.133,90.153,90.213,90.206,90.192,90.215,90.292,90.373,90.414,90.386,90.454,90.44,90.083]
[]
[]
["Neutral 8"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},56.374706,58.551815,47.843143,81.044615,-0.200032,0.528512]
["Gretag spec.","DE spec Lab:",{"v":0.007077537142255113,"d":"0.007"},null,56.376189,58.55184,47.837924,81.044629,-0.196426,0.534602,13.954,21.246,35.511,50.422,56.381,57.61,57.957,58.247,58.504,58.508,58.485,58.553,58.604,58.534,58.495,58.424,58.525,58.543,58.406,58.684,58.698,58.734,58.717,58.677,58.569,58.411,58.35,58.235,58.216,58.028,57.822,57.568,57.548,57.455,57.375,57.293]
[null,"DE to Gretag XYZ.:",{"v":0.013491462930314698,"d":"0.013"}]
["Argyll","DE to Gretag spec.:",{"v":0.019338723225692925,"d":"0.019"},null,56.386909,58.564608,47.863518,81.051683,-0.20033,0.5170239999999999,13.967,21.295,35.548,50.561,56.432,57.646,58.011,58.3,58.502,58.529,58.509,58.561,58.614,58.541,58.504,58.435,58.543,58.568,58.429,58.704,58.716,58.75,58.707,58.666,58.57,58.411,58.358,58.24,58.204,58.014,57.836,57.584,57.584,57.468,57.391,57.325]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.03008920188373225,"d":"0.030"},null,56.395433,58.567604,47.88889,81.053338,-0.18639599999999998,0.490407,12.806000000000001,13.09,14.519,16.688,19.906,23.991,29.205,35.089,41.18,46.792,51.228,54.147,55.807,56.716,57.194,57.449,57.659,57.871,57.918,57.989,58.092,58.249,58.378,58.454,58.506,58.538,58.532,58.53,58.547,58.528,58.515,58.543,58.521,58.529,58.564,58.59,58.592,58.604,58.595,58.593,58.579,58.526,58.498,58.543,58.481,58.401,58.403,58.48,58.546,58.586,58.583,58.557,58.539,58.507,58.466,58.431,58.473,58.589,58.698,58.709,58.69,58.728,58.767,58.784,58.778,58.758,58.731,58.709,58.672,58.678,58.678,58.668,58.631,58.576,58.523,58.472,58.417,58.385,58.374,58.353,58.29,58.2,58.237,58.384,58.372,58.217,58.109,58.02,57.98,57.932,57.899,57.835,57.721,57.617,57.589,57.597,57.601,57.596,57.529,57.457,57.432,57.438,57.43,57.434,57.378,57.338,57.33,57.35,57.305,57.283,57.327,57.309,57.047]
[]
[]
["Neutral 6.5"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},34.236649,35.499397,28.94685,66.135641,0.028027,0.54469]
["Gretag spec.","DE spec Lab:",{"v":0.006144493225645209,"d":"0.006"},null,34.23753,35.499418,28.943556,66.135657,0.030997,0.550069,12.416,17.942,26.257,32.441,34.236,34.716,35.023,35.257,35.378,35.345,35.263,35.281,35.323,35.305,35.318,35.314,35.41,35.435,35.409,35.663,35.762,35.818,35.797,35.738,35.601,35.441,35.356,35.19,35.101,34.894,34.697,34.475,34.416,34.248,34.15,34.033]
[null,"DE to Gretag XYZ.:",{"v":0.010800540773502194,"d":"0.011"}]
["Argyll","DE to Gretag spec.:",{"v":0.016486951082598966,"d":"0.016"},null,34.230491,35.495464,28.947573,66.132607,0.019877,0.538285,12.219,17.951,26.197,32.365,34.211,34.73,35.016,35.268,35.389,35.347,35.281,35.292,35.341,35.312,35.323,35.316,35.401,35.438,35.408,35.649,35.743,35.811,35.788,35.722,35.594,35.434,35.337,35.186,35.1,34.893,34.707,34.483,34.397,34.259,34.161,34.032]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.014035813371517864,"d":"0.014"},null,34.242911,35.507139,28.963285,66.141611,0.023881,0.52829,10.736,11.71,12.892,14.792,17.122,19.956,23.189,26.304,29.035,31.192,32.608,33.516,34.044,34.272,34.466,34.595,34.713,34.856,34.887,34.998,35.102,35.207,35.317,35.41,35.411,35.396,35.443,35.37,35.348,35.321,35.33,35.335,35.305,35.286,35.304,35.325,35.32,35.337,35.358,35.367,35.334,35.313,35.322,35.35,35.333,35.275,35.292,35.361,35.418,35.451,35.441,35.421,35.414,35.422,35.423,35.409,35.462,35.579,35.664,35.699,35.723,35.776,35.818,35.833,35.837,35.841,35.82,35.803,35.768,35.741,35.739,35.707,35.653,35.596,35.558,35.508,35.441,35.391,35.368,35.34,35.272,35.199,35.171,35.253,35.237,35.097,35.009,34.93,34.843,34.828,34.786,34.717,34.615,34.537,34.488,34.475,34.448,34.42,34.362,34.281,34.225,34.226,34.214,34.161,34.116,34.088,34.058,34.017,33.968,33.9,33.886,33.845,33.674]
[]
[]
["Neutral 5"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},18.561424,19.272818,16.073835,51.004874,-0.11104599999999999,-0.423994]
["Gretag spec.","DE spec Lab:",{"v":0.004988278761256232,"d":"0.005"},null,18.561886,19.272823,16.072014,51.00488,-0.108681,-0.419602,10.034,13.202,16.805,18.803,19.26,19.422,19.497,19.616,19.602,19.529,19.449,19.402,19.386,19.328,19.283,19.248,19.275,19.252,19.21,19.331,19.358,19.36,19.319,19.269,19.171,19.041,18.94,18.822,18.743,18.594,18.467,18.325,18.252,18.155,18.07,17.989]
[null,"DE to Gretag XYZ.:",{"v":0.029800493301284287,"d":"0.030"}]
["Argyll","DE to Gretag spec.:",{"v":0.027692893601065274,"d":"0.028"},null,18.573945,19.287826,16.077147,51.022262,-0.121097,-0.401978,9.995,13.253,16.786,18.79,19.251,19.422,19.519,19.602,19.605,19.54,19.47,19.422,19.399,19.342,19.302,19.27,19.291,19.269,19.226,19.333,19.373,19.379,19.341,19.282,19.17,19.05,18.959,18.841,18.755,18.616,18.479,18.33,18.277,18.167,18.098,17.998]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.029029496637040648,"d":"0.029"},null,18.572649,19.28325,16.083139,51.016961,-0.104962,-0.425521,9.189,9.646,10.511,11.555,12.913,14.322,15.7,16.916,17.858,18.529,18.921,19.092,19.222,19.314,19.327,19.353,19.418,19.464,19.476,19.513,19.547,19.589,19.644,19.658,19.639,19.617,19.597,19.563,19.535,19.519,19.485,19.472,19.441,19.407,19.427,19.402,19.405,19.394,19.379,19.372,19.345,19.324,19.303,19.315,19.274,19.239,19.239,19.275,19.304,19.3,19.291,19.262,19.247,19.246,19.227,19.212,19.237,19.291,19.333,19.344,19.342,19.37,19.397,19.394,19.382,19.386,19.364,19.336,19.304,19.293,19.279,19.256,19.206,19.167,19.131,19.086,19.058,19.014,18.974,18.953,18.897,18.858,18.832,18.863,18.84,18.76,18.697,18.642,18.607,18.576,18.53,18.49,18.429,18.356,18.331,18.324,18.298,18.285,18.238,18.18,18.155,18.144,18.127,18.104,18.07,18.014,17.994,17.978,17.937,17.908,17.925,17.909,17.803]
[]
[]
["Neutral 3.5"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},8.240786,8.548365,6.982274,35.099777,-0.013786,0.28945000000000004]
["Gretag spec.","DE spec Lab:",{"v":0.0038947540358795067,"d":"0.004"},null,8.240991,8.548369,6.981453,35.099786,-0.011994,0.292908,6.477,7.256,7.8870000000000005,8.16,8.292,8.375,8.493,8.521,8.539,8.46,8.417,8.409,8.449,8.469,8.5,8.518,8.544,8.563,8.548,8.608,8.635,8.635,8.621,8.588,8.537,8.487,8.453,8.392,8.355,8.311,8.259,8.194,8.178,8.139,8.106,8.085]
[null,"DE to Gretag XYZ.:",{"v":0.022510474539642476,"d":"0.023"}]
["Argyll","DE to Gretag spec.:",{"v":0.02422763407763812,"d":"0.024"},null,8.245598,8.555407,6.988123,35.113805,-0.031388,0.289124,6.409,7.204,7.934,8.225,8.265,8.376,8.48,8.555,8.535,8.475,8.427,8.432,8.458,8.482,8.511,8.526,8.558,8.564,8.557,8.614,8.634,8.646,8.627,8.589,8.542,8.489,8.447,8.398,8.365,8.304,8.253,8.191,8.173,8.141,8.111,8.075]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.014270213242974228,"d":"0.014"},null,8.250891,8.559738,6.992986,35.122429,-0.021439,0.283621,6.357,6.319,6.556,6.873,7.172,7.552,7.747,7.922,8.026,8.104,8.223,8.241,8.23,8.295,8.318,8.327,8.386,8.411,8.438,8.466,8.509,8.553,8.568,8.585,8.56,8.546,8.52,8.506,8.485,8.464,8.449,8.434,8.427,8.427,8.447,8.441,8.444,8.454,8.472,8.484,8.495,8.5,8.498,8.513,8.516,8.514,8.518,8.539,8.551,8.57,8.57,8.569,8.564,8.566,8.565,8.554,8.575,8.602,8.621,8.633,8.625,8.633,8.659,8.661,8.65,8.645,8.638,8.623,8.621,8.606,8.592,8.588,8.571,8.547,8.541,8.526,8.496,8.473,8.463,8.458,8.429,8.418,8.406,8.414,8.389,8.354,8.333,8.312,8.291,8.289,8.288,8.26,8.248,8.223,8.208,8.183,8.179,8.19,8.156,8.137,8.135,8.132,8.132,8.131,8.109,8.094,8.097,8.07,8.056,8.062,8.053,8.021,7.994]
[]
[]
["Black"]
[null,null,"DE's","Reading","X","Y","Z","L*","a*","b*",{"v":380,"d":"380.000"},{"v":390,"d":"390.000"},{"v":400,"d":"400.000"},{"v":410,"d":"410.000"},{"v":420,"d":"420.000"},{"v":430,"d":"430.000"},{"v":440,"d":"440.000"},{"v":450,"d":"450.000"},{"v":460,"d":"460.000"},{"v":470,"d":"470.000"},{"v":480,"d":"480.000"},{"v":490,"d":"490.000"},{"v":500,"d":"500.000"},{"v":510,"d":"510.000"},{"v":520,"d":"520.000"},{"v":530,"d":"530.000"},{"v":540,"d":"540.000"},{"v":550,"d":"550.000"},{"v":560,"d":"560.000"},{"v":570,"d":"570.000"},{"v":580,"d":"580.000"},{"v":590,"d":"590.000"},{"v":600,"d":"600.000"},{"v":610,"d":"610.000"},{"v":620,"d":"620.000"},{"v":630,"d":"630.000"},{"v":640,"d":"640.000"},{"v":650,"d":"650.000"},{"v":660,"d":"660.000"},{"v":670,"d":"670.000"},{"v":680,"d":"680.000"},{"v":690,"d":"690.000"},{"v":700,"d":"700.000"},{"v":710,"d":"710.000"},{"v":720,"d":"720.000"},{"v":730,"d":"730.000"}]
["Gretag XYZ",{"$nulls":3},3.029839,3.119925,2.5617900000000002,20.517923,0.375951,0.096673]
["Gretag spec.","DE spec Lab:",{"v":0.002734396094204371,"d":"0.003"},null,3.029918,3.119929,2.5615,20.517935,0.377274,0.099066,3.132,3.166,3.12,3.108,3.143,3.098,3.1,3.113,3.108,3.087,3.105,3.104,3.101,3.099,3.096,3.09,3.104,3.1,3.1,3.116,3.123,3.138,3.143,3.155,3.169,3.182,3.21,3.232,3.2640000000000002,3.274,3.289,3.312,3.342,3.377,3.391,3.41]
[null,"DE to Gretag XYZ.:",{"v":0.04315382831221276,"d":"0.043"}]
["Argyll","DE to Gretag spec.:",{"v":0.043663588560719925,"d":"0.044"},null,3.038807,3.12785,2.570782,20.548816,0.398308,0.076473,3,3.115,3.174,3.147,3.148,3.137,3.128,3.115,3.107,3.103,3.098,3.099,3.111,3.116,3.106,3.099,3.107,3.106,3.106,3.122,3.1310000000000002,3.144,3.158,3.17,3.179,3.191,3.213,3.23,3.263,3.28,3.296,3.32,3.354,3.382,3.403,3.437]
[{"$nulls":3},"Reading","X","Y","Z","L*","a*","b*",{"v":376.666667,"d":"376.667"},{"v":380,"d":"380.000"},{"v":383.333333,"d":"383.333"},{"v":386.666667,"d":"386.667"},{"v":390,"d":"390.000"},{"v":393.333333,"d":"393.333"},{"v":396.666667,"d":"396.667"},{"v":400,"d":"400.000"},{"v":403.333333,"d":"403.333"},{"v":406.666667,"d":"406.667"},{"v":410,"d":"410.000"},{"v":413.333333,"d":"413.333"},{"v":416.666667,"d":"416.667"},{"v":420,"d":"420.000"},{"v":423.333333,"d":"423.333"},{"v":426.666667,"d":"426.667"},{"v":430,"d":"430.000"},{"v":433.333333,"d":"433.333"},{"v":436.666667,"d":"436.667"},{"v":440,"d":"440.000"},{"v":443.333333,"d":"443.333"},{"v":446.666667,"d":"446.667"},{"v":450,"d":"450.000"},{"v":453.333333,"d":"453.333"},{"v":456.666667,"d":"456.667"},{"v":460,"d":"460.000"},{"v":463.333333,"d":"463.333"},{"v":466.666667,"d":"466.667"},{"v":470,"d":"470.000"},{"v":473.333333,"d":"473.333"},{"v":476.666667,"d":"476.667"},{"v":480,"d":"480.000"},{"v":483.333333,"d":"483.333"},{"v":486.666667,"d":"486.667"},{"v":490,"d":"490.000"},{"v":493.333333,"d":"493.333"},{"v":496.666667,"d":"496.667"},{"v":500,"d":"500.000"},{"v":503.333333,"d":"503.333"},{"v":506.666667,"d":"506.667"},{"v":510,"d":"510.000"},{"v":513.333333,"d":"513.333"},{"v":516.666667,"d":"516.667"},{"v":520,"d":"520.000"},{"v":523.333333,"d":"523.333"},{"v":526.666667,"d":"526.667"},{"v":530,"d":"530.000"},{"v":533.333333,"d":"533.333"},{"v":536.666667,"d":"536.667"},{"v":540,"d":"540.000"},{"v":543.333333,"d":"543.333"},{"v":546.666667,"d":"546.667"},{"v":550,"d":"550.000"},{"v":553.333333,"d":"553.333"},{"v":556.666667,"d":"556.667"},{"v":560,"d":"560.000"},{"v":563.333333,"d":"563.333"},{"v":566.666667,"d":"566.667"},{"v":570,"d":"570.000"},{"v":573.333333,"d":"573.333"},{"v":576.666667,"d":"576.667"},{"v":580,"d":"580.000"},{"v":583.333333,"d":"583.333"},{"v":586.666667,"d":"586.667"},{"v":590,"d":"590.000"},{"v":593.333333,"d":"593.333"},{"v":596.666667,"d":"596.667"},{"v":600,"d":"600.000"},{"v":603.333333,"d":"603.333"},{"v":606.666667,"d":"606.667"},{"v":610,"d":"610.000"},{"v":613.333333,"d":"613.333"},{"v":616.666667,"d":"616.667"},{"v":620,"d":"620.000"},{"v":623.333333,"d":"623.333"},{"v":626.666667,"d":"626.667"},{"v":630,"d":"630.000"},{"v":633.333333,"d":"633.333"},{"v":636.666667,"d":"636.667"},{"v":640,"d":"640.000"},{"v":643.333333,"d":"643.333"},{"v":646.666667,"d":"646.667"},{"v":650,"d":"650.000"},{"v":653.333333,"d":"653.333"},{"v":656.666667,"d":"656.667"},{"v":660,"d":"660.000"},{"v":663.333333,"d":"663.333"},{"v":666.666667,"d":"666.667"},{"v":670,"d":"670.000"},{"v":673.333333,"d":"673.333"},{"v":676.666667,"d":"676.667"},{"v":680,"d":"680.000"},{"v":683.333333,"d":"683.333"},{"v":686.666667,"d":"686.667"},{"v":690,"d":"690.000"},{"v":693.333333,"d":"693.333"},{"v":696.666667,"d":"696.667"},{"v":700,"d":"700.000"},{"v":703.333333,"d":"703.333"},{"v":706.666667,"d":"706.667"},{"v":710,"d":"710.000"},{"v":713.333333,"d":"713.333"},{"v":716.666667,"d":"716.667"},{"v":720,"d":"720.000"},{"v":723.333333,"d":"723.333"},{"v":726.666667,"d":"726.667"},{"v":730,"d":"730.000"},{"v":733.333333,"d":"733.333"},{"v":736.666667,"d":"736.667"},{"v":740,"d":"740.000"},{"v":743.333333,"d":"743.333"},{"v":746.666667,"d":"746.667"},{"v":750,"d":"750.000"}]
["Argyll HiRes","DE to Argyll Norm:",{"v":0.06467849621010087,"d":"0.065"},null,3.041971,3.129943,2.579997,20.556964,0.417996,0.015404999999999999,2.89,3.081,3.047,3.105,3.11,3.08,3.173,3.214,3.234,3.27,3.236,3.177,3.191,3.208,3.174,3.15,3.155,3.169,3.136,3.133,3.14,3.115,3.124,3.125,3.12,3.132,3.11,3.099,3.11,3.107,3.101,3.1,3.108,3.105,3.108,3.113,3.105,3.125,3.114,3.118,3.125,3.11,3.108,3.11,3.111,3.099,3.097,3.106,3.101,3.105,3.109,3.111,3.107,3.113,3.107,3.098,3.108,3.119,3.126,3.124,3.128,3.1390000000000002,3.145,3.143,3.145,3.147,3.159,3.168,3.16,3.161,3.169,3.176,3.174,3.177,3.185,3.19,3.189,3.205,3.215,3.216,3.216,3.222,3.237,3.247,3.2560000000000002,3.262,3.276,3.267,3.275,3.3,3.3,3.292,3.297,3.314,3.308,3.335,3.348,3.341,3.355,3.359,3.385,3.4,3.418,3.427,3.422,3.427,3.439,3.437,3.463,3.453,3.459,3.477,3.439]